## Slide 1: Artifacts, ghosts, and phantoms: An archaeological tour of MR Imaging
Jason Stockmann
Why N How
Oct. 31, 2019

## Slide 2
Interesting…
What kind of artifact is this…?
Where does it come from…?
How was it made…?
MRI Physicist
MRI physicist =
Professional artifact hunter

## Slide 3: Acknowledgements
Most of the slides borrowed from the lecture notes and presentations of Lawrence Wald, Todd Constable, Jon Polimeni, and Doug Noll

## Slide 4: Why does MRI have such a wide variety of image artifacts?
MRI requires seamless integration and interplay of magnet, gradient coil system, RF system, and pulse sequence computer, each of which is a complicated subsystem.
Human body also moves and interacts with RF and B0 magnetic fields  both static and dynamic (physiological) effects
Very old figure…!

## Slide 5
Learn how to pinpoint where the source of your artifact is hidden in the “k-space domain”… You need to know where to look!
K-space city

## Slide 6: Understanding artifacts
Data is acquired in k-space domain  2D or 3D Fourier transform of the image
Learn to think about what’s going wrong with data collection in k-space  use properties of Fourier transform to see what effect these errors will create in image domain

## Slide 7: Fourier Transform
Shows you the spatial frequencies that are present in the image (or audio signal, etc.).
Steps:
  (1) Multiply input function (image) by spatial harmonic (sinusoid) with a given frequency
  (2) Integrate over the whole space (1D, 2D, or 3D)  Gives you one point in k-space
  (3) Change spatial frequency (k) and repeat steps (1) and (2) until you have populated all of k-space
One k-space point tells you how much of the corresponding spatial harmonic is needed to build the image. When you add all the harmonics together, the complex-valued functions interfere constructively and destructively in a way that forms the image features.
Jean-Baptiste Joseph Fourier
(1768-1830)
Measured signal is Fourier integral of the projection image!
1D Fourier transform along x
M0 is the object
kx is spatial frequency (k-space coordinate)
In practice we use the discretized version of this formula. Number of k-space points depends on size of image grid

## Slide 8
ky
1 / Δx
1 / FOVx
FOVx = matrix * Resx
Fourier transform
You’ve measured:
intensity at a spatial frequency...
kx
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 9
Planning your image…
FOVx = matrix * Resx
Figure out what you want
1 / FOVx
ky
1 / Δx
kx
Figure out your kspace params
Figure out your gradient params
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 10
Planning your image…
ky
Nx
kx
Data lands in integer indexed matrix
Ny
Ta Da!
Nx
Ny
Matlab fft2
(Unitary, reversible operator)
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH
Brains?

## Slide 11: Gradient: Basics of image encoding
Uniform magnet
Field from gradient coils
Total field
Bo Gx x Bo + Gx x
x
z
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 12
Recasting the signal equation
Change variables:
Measured signal is Fourier integral of the projection image!
Science/engineering principle #1: cast your problem into the form of a problem with a known solution…
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 13
ky
kx
RF
Gz
Gx
Gy
Sample
One shot per readout line…
TE
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 14: Image encoding strategies: Echo Planar Imaging
ky
kx
RF
Gz
Gx
Gy
Sample
All lines in one shot…
esp
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 15
Top 5 features of Fourier transform
1) Information content is equal in the two domains.
  (inverse FT recovers the periodic function)
256 x 256 complex #s
256 x 256 complex #s
F(kx,ky)
f(x,y)
FT-1
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 16
Top 5 features of Fourier transform
1) Information content is equal in the two domains.
  (inverse FT recovers the periodic function)
256 x 256 complex #s
256 x 256 complex #s
F(kx,ky)
f(x,y)
FT-1
Δkx= 1/FOVx
2kxmax =1/Δx
FOVx
FOVy
Δx= pixel size
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 17
Top 6 features of Fourier transform
  If you make FOV too small (Δk to big), you will get aliasing (image wrap).
FT-1
kx
ky
Only green samples
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 18
Top 5 features of Fourier transform
2) Even though a real object has 0 phase, phase is incredibly important in k-space.
256 x 256 complex #s
256 x 256 complex #s
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 19
Where is the majority of the information?
magnitude or phase?
magnitude
phase
magnitude
phase
Larry
Christin
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH
K-space:

## Slide 20
1) Larry (mag. dominate) Christin
2) Christin (phase dominate) Larry
3) Mutant monster (both critical) Mutant monster
Where is the majority of the information?
magnitude or phase?
magnitude
phase
magnitude
phase
Larry
Christin
??
??
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 21
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH
Where is the majority of the information?
magnitude or phase?
magnitude
phase
magnitude
phase
Larry
Christin

## Slide 22
Top 5 features of Fourier transform
3) Narrow in one domain is wide in the other.
Time/frequency domains
uo
u
t
Mxy
uo
u
t
Mxy
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 23
K-space
Image domain
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 24
Top 5 features of Fourier transform
3) Narrow in one domain is wide in the other.
Kspace/space domains
One “white pixel” in kspace from a electric spark
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 25
Single Spike in Raw Data: Only Location and Amp of Spike is Changing
Slide courtesy of Todd Constable, Yale Univ.

## Slide 26
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 27: Filtering and k-space
x =
x =
edge enhancement
edge blurring
Slide courtesy of Todd Constable, Yale Univ.

## Slide 28
Top 5 features of Fourier transform
4) Shift Theorem
  Translation in space is a linear phase roll in kspace
Let F(k) = FT{ f(x) }
FT{ f(x-a) } = F(k) e-2πjak
mag
phase
FT
+
FT
Adapted from slide provided by Lawrence Wald, A. A. Martinos Center, MGH

## Slide 29
Top 5 features of Fourier transform
5) Convolution Theorem
  Multiplication of 2 different kspaces is a FT of a convolution in space
Let: F(k) = FT{ f(x) } R(k) = FT{ r(x) }
Convolution:
FT{ f(x) r(x)} = F(k) R(k)
*
“FT of a convolution
Is the product of the FTs”
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH
….leads to Gibbs ringing

## Slide 30: Artifacts (a partial list)
Hardware/Acquisition
Human body, physiology, implants
Gradient nonlinearity
Magnet drift
Eddy currents
RF (zipper)
RF interference (zipper)
Truncation
Aliasing
Ghosting
Sequence-specific (streaking in radial trajectories, blurring in spiral)
Susceptibility (rare)
Receiver dynamic range clipping (rare)
Susceptibility: distortion
Susceptibility: dropout/dephasing & implants
RF field inhomogeneity (coil bias)
Motion
Flow
Chemical shift
Subject-specific and generally harder to fix!

## Slide 31
asymmetry in z leads to some gradient nonlinearity
geometric distortion
intensity bias
spatially nonuniform voxel size
correctable
Slide courtesy of Jon Polimeni, MGH

## Slide 32: gradient nonlinearity distortion
distortion is same for any object, pulse sequence—
depends only on gradient coil
original, acquired
online correction
Slide courtesy of Jon Polimeni, MGH

## Slide 33: “ball-grid” validation phantom
http://www.birncommunity.org/resources/supplements/brain-morphometry-multi-site-studies/brain-morphometry-mri-phantom/
250 mm length × 220 mm dia.
water-filled beads, 3 mm dia.
spaced at 10 +/− 0.05 mm
(too large for most RF coils)
AC84 head gradient coil
Slide courtesy of Sebastian Littin, Univ. of Freiburg

## Slide 34: “ball-grid” validation phantom
http://www.birncommunity.org/resources/supplements/brain-morphometry-multi-site-studies/brain-morphometry-mri-phantom/
250 mm length × 220 mm dia.
water-filled beads, 3 mm dia.
spaced at 10 +/− 0.05 mm
(too large for most RF coils)
AC84 head gradient coil
measuring resulting spacing to validate unwarping
Slide courtesy of Sebastian Littin, Univ. of Freiburg

## Slide 35: surface-based grad. nonlin. correction
distorted
corrected
avoid blurring by deforming surfaces
resampling requires interpolation,
blurs image
Slide courtesy of Jon Polimeni, MGH

## Slide 36
7T head gradient in-fold artifact
TEM 3
end-capped
27cm dia
14cm long
Shoulder/Brain= 33%, 21%
Body gradients
Head gradients
Coil picks up signal from shoulders…

## Slide 37: Zipper Artifacts
RF interference  usually shows up as lines along phase-encode direction
Room not properly shielded
RF source may be in room - electronic devices
Slide courtesy of Todd Constable, Yale Univ.

## Slide 38: Truncation Artifacts: Consequence of Convolution rule
Echo signal
Fourier transform (image)
Slide courtesy of Todd Constable, Yale Univ.

## Slide 39: Truncation Artifacts (Gibbs ringing)
Truncation artifacts can mimic disease - case of true syringomyelia above left, treated and resolved on the right. Looks like a truncation artifact but in this case is not.
Mayo Clin Proc. 2002;77:291-294 © 2002 Mayo Foundation for Medical Education and Research
Slide courtesy of Todd Constable, Yale Univ.

## Slide 40: Eddy currents
Switching magnetic field creates unwanted currents in other nearby conductors in the bore (Lenz’s law)
Resists the desired change in gradient field amplitude

## Slide 41: Distortions: Eddy currents
Cause: Fast switching of diffusion-encoding gradients induces eddy currents in conducting components
The shifts are direction-dependent, i.e., different for each DW image
Result: Geometric shifts/distortions
Figure courtesy of Douglas Noll, Univ. Michigan
(readout)
(phase encode)

## Slide 42: Aliasing: FOV wrap
K-space samples are too far apart (results in FOV too small)
Image is wrapped around
FOVy=40cm
FOVy=20cm
If resolution is to be maintained between above 2 acquisitions, imaging time will double
In many sequences, you can save time if you skip k-space lines and allow aliasing…
Parallel imaging using multiple RF receive coils “unwraps” aliased image (“iPAT”, “GRAPPA”)
Slide courtesy of Todd Constable, Yale Univ.

## Slide 43
Spiking  actively powered components, usually RF coil (PIN diode lines) or gradient coil

## Slide 44
Spiking in EPI time series  comes and goes over time
Usually a problem with “active” powered components like gradient coil or RF coil  small arc of electric current generates spike

## Slide 45
Adjust brightness and contrast…
On Siemens scanner, click with right mouse button and choose “Select series”
Then click with middle mouse button and drag down and to the left
Then the spikes emerge…
13th century Welsh medieval manuscript of Arthurian legends under UV light

## Slide 46: Ghosts!
Bipolar readout gradient  Even and odd lines of k-space are slightly different
Creates image shifted by half a field of view (“N/2 ghost”)
Phase in k-space causes translation of some of the energy in image doman
“Navigators” used to fix ghost
But EPI is very sensitive to hardware instability that can gause ghosting
Le Bihan, JMRI 2006
Figure courtesy of Douglas Noll, Univ. Michigan

## Slide 47: Ghost flickering from hardware instability  not the same as “spiking”

## Slide 48: Ghost flickering from hardware instability  not the same as “spiking”
Insidious, evil artifact, especially at 7T…. Can bias EPI time series analysis
Possible causes:
RF power amplifier
RF Tx or Rx coil
Too few references lines for training GRAPPA kernel
GRAPPA factor is too high for the coil array  GRAPPA starts to fail
Using a “forbidden” echo spacing  gradient coil acoustic resonance!
Subject motion during ACS lines for GRAPPA

## Slide 49: Artifacts
Hardware/Acquisition
Human body, physiology, implants
Gradient nonlinearity
Magnet drift
Eddy currents
RF (zipper)
RF interference (zipper)
Truncation
Aliasing
Ghosting
Sequence-specific (streaking in radial trajectories, blurring in spiral)
Susceptibility (rare)
Receiver dynamic range clipping (rare)
Susceptibility: distortion
Susceptibility: dropout/dephasing & implants
RF field inhomogeneity (coil bias)
Motion
Flow
Chemical shift
Subject-specific and typically harder to fix!!!

## Slide 50: Source of in vivo ΔB0 inhomogeneity
Air cavities
water
Empty magnet is shimmed very well accurately
Air-tissue interfaces in body perturb B0 field by up to ~1ppm
  Tissue: diamagnetic
  Air: paramagnetic
Static ΔB0:
  Sinus and oral cavities
  Ear canals
Dynamic ΔB0:
  Chest motion in respiration
  Other physiological processes
7.00001
6.99999
7.00000
Source: NIH National Cancer Institute.
http://www.cancer.gov
Slide courtesy of Jiazheng Zhou

## Slide 51
200
-200
0
Hz
Static ΔB0 shimming at UHF
7T ΔB0 field map shimmed to 2nd order
Susceptibility-contrast (T2*) methods such as blood-oxygen level dependent (BOLD) fMRI and high-res. Susceptibility-weighted imaging (SWI)
Interested in microscopic T2* of tissue
Confounded by macroscopic T2* due to ΔB0
  grows with B0
Scanner tries to “shim” out these field perturbations, but only up to 2nd order spatial components can be removed…
𝜎B0GLOBAL = 52 Hz
Stockmann JP, NeuroImage, 2017

## Slide 52
Images courtesy of Jon Polimeni, MGH
7T gradient echo mag. and phase images, 300 μm in-plane, cardiac-gated

## Slide 53: Impact of static ΔB0 on macroscopic T2* at 7T
Hz
ΔB0 field maps
T2* maps
1.5×1.5×2 mm
1.1 mm iso.
T2-weighted anatomical reference
T2*
(ms)
+200
-200
0
0
30
15
Stockmann JP, NeuroImage, 2017

## Slide 54
Through-slice dropout in 7T gradient-echo EPI slices
6 mm slice
4 mm
2 mm
0.75 mm
1 mm
Intravoxel dephasing scales with voxel size
Also gets worse with longer TEs
Images courtesy of Jon Polimeni, MGH

## Slide 55: Metal implants
Severely short T2* near metal in the body
  Dephasing / signal loss in gradient echo
  Warping / distortion
  RF excitation and refocusing pulses don’t work properly
Metal in scalp
Dental metal
Hip implant
Slide courtesy of Todd Constable, Yale Univ.

## Slide 56: Metal Implants
Spin echo
Gradient echo
Ferromagnetic Dental Work & a Hair Band, Piercings, Tattoos, Colored Contact Lens
Slide courtesy of Todd Constable, Yale Univ.

## Slide 57: Geometric distortion in EPI
Depends on B0 off-resonance frequency and the sequence’s phase encoding bandwidth
Linear phase ramp evolves during echo train (along k-space phase encode direction)
  Causes voxel shift along phase encode direction
Phase encode bandwidth is 1/echo_spacing
So for 1ms echo spacing, and image with 100 pixels in PE direction, 100 Hz off-resonance gives you a 10 pixel shift!
Bigger off-resonance  larger shifts
GRAPPA (Ry) reduces distortion! Reduces effective echo spacing
200
-200
0
Hz
Reference: Wald LL, The future of acquisition speed, coverage, sensitivity, and resolution. Neuroimage, 2012.

## Slide 58
5) Shift Theorem
  Translation in space is a linear phase roll in kspace
Let F(k) = FT{ f(x) }
FT{ f(x-a) } = F(k) e-2πjak
mag
phase
FT
+
FT
Adapted from slide provided by Lawrence Wald, A. A. Martinos Center, MGH
Reminder!

## Slide 59
Less common: Increase gradient strength and slew rate
EPI
3T images
EPI
7T images
ΔB0 map
ΔB0 map
Ref.
Ref.
Hz
A-P
P-A
A-P
P-A
= 51.9 Hz
= 21.4 Hz
Geometric distortion in EPI
TE=30ms
TE=23ms
Most common mitigation:
parallel imaging2 acceleration  reduce effective echo spacing
Post-processing unwarping algorithms like TOPUP1 help, but are not perfect
1Andersson JLR, Neuroimage, 2003
2Griswold M, Mag. Res. Med., 2002
Less common: Increase gradient strength and slew rate
Stockmann JP, NeuroImage, 2017

## Slide 60: EPI with alternating slice-encoding gradients: over-easy
[Błażejewska et al., 2017 ISMRM]
7T
1.1 mm iso.
axial acquisition
A→P phase encode
alternating slice-select gradient polarity
“over-easy” EPI
Slide courtesy of Jon Poliemeni

## Slide 61: Dynamic ΔB0 caused by physiology
Ghosting and ringing in structural images
  K-space phase modulation due to B0 fluctuations  ghosts
  phase stabilization navigators help…
Stability of EPI time-series functional images
  Respiration explains ~65% of variance in phase time-series
Regress out in post-processing
PhaseEPI(x,y,z,t)
Figure courtesy of Marta Bianciardi, MGH

## Slide 62
Spatiotemporal B0 fluctuations are worse at 7T, especially in deep brain
Feedback-controlled dynamic B0 shim updating could help….
Gross et al., ISMRM 2016
Emerging method: Use real-time field monitoring and B0 shimming to reduce artifacts in structural images

## Slide 63: tSNR discontinuity in accel. EPI: chest motion during respiration
free breathing during ACS acquisition
breath-hold during ACS acquisition
FLEET prevents phase errors between shots caused by dynamic B0 changes
conventional
tSNR
conventional
tSNR
FLEET (αi = 5°)
tSNR
FLEET (αi = 5°)
tSNR
conventional
conventional
std.
std.
Slide courtesy of Jon Polimeni, MGH

## Slide 64
vulnerability to motion assessed with a mechanical, anthropomorphic head phantom rocking 5°every 15 s
effects of motion: conventional ACS
Slide courtesy of
Jon Polimeni, MGH

## Slide 65
effects of motion: FLEET ACS
vulnerability to motion assessed with a mechanical, anthropomorphic head phantom rocking 5°every 15 s
Slide courtesy of
Jon Polimeni, MGH

## Slide 66
Surface Coil B1- Sensitivity Profile Effects
4 surface coils and their sensitivity profiles resultant magnitude images
coil 1 coil 2 image 1 image 2
coil 3 coil 4 image 3 image 4
Combined is brighter where Rx coils are more sensitive
Slide courtesy of Todd Constable, Yale Univ.

## Slide 67: RF B1+ inhomogeneity at ultra-high field (7T)
Interaction of transmit RF field with body tissue (dielectric permittivity and conductivity)
Dark shading in images, loss of contrast
Ugurbil K, NeuroImage, 2017.
B1+ field amplitude
7T brain images

## Slide 68: Motion Artifacts
Motion causes to phase errors in the data which leads to mis-registration in the phase encode (y) direction
periodic motion leads to ghosting (shifted images that can constructively or destructively interfere with the primary image)
Remedies:
  correct for motion induced phase errors
  fast imaging to freeze motion
  follow motion
Slide courtesy of Todd Constable, Yale Univ.

## Slide 69: 1D Example: Motion Present
consider a 1D image consisting of a box
  periodic motion
data domain (time/k-space)
image domain
Fourier Transform
original
k-space data
Fourier Transform
*
=
ghost images
.
modulation from periodic motion
Slide courtesy of Todd Constable, Yale Univ.

## Slide 70: 1D Example: Motion Present
consider a 1D image consisting of a box
  periodic motion
modulation of original data by periodic phase errors leads to discrete ghost images
*
=
.
FT
periodic motion
Slide courtesy of Todd Constable, Yale Univ.
original
k-space data

## Slide 71: 1D Example: Motion Present
consider a 1D image consisting of a box
  periodic motion
ghosts separated in this example
usually overlap in practice
*
=
.
FT
periodic motion
Slide courtesy of Todd Constable, Yale Univ.
original
k-space data

## Slide 72: Motion Artifacts in 2D: Conventional Data Acquisition
2DFT
data acquired 1 line at a time with a delay = TR between each line
Slide courtesy of Todd Constable, Yale Univ.

## Slide 73: Motion Artifacts in 2D: Conventional Data Acquisition
periodic motion introducing a phase error in the phase encode (vertical) direction results in misplacement of tissue in phase encode direction
2DFT
Slide courtesy of Todd Constable, Yale Univ.

## Slide 74: Motion Artifacts in 2D: Conventional Data Acquisition
phase errors may occur over the datasets for only certain voxels - not all voxels
(e.g. only the moving ones)
2DFT
Slide courtesy of Todd Constable, Yale Univ.

## Slide 75: Motion Artifacts in 2D: Different severity/extent of artifacts
In this case motion included entire object
Often limited to anterior wall, or specific organs/structures
Motion artifact primarily from beating heart
Motion artifact from flowing blood in vessel
periodic due to pulsatile nature of flow
Slide courtesy of Todd Constable, Yale Univ.

## Slide 76: Head motion
No motion
No motion; with MoCo
With motion; no MoCo
With motion; no MoCo
d
With motion and MoCo
3D acquisition example:
T1-weighted MPRAGE

## Slide 77: Fat and Water Resonances Move in and out of phase over time
Water precesses approx. 3.5ppm faster than fat
Slide courtesy of Todd Constable, Yale Univ.

## Slide 78: Chemical Shift can be Exploited to Produce Fat (left) or Water (right) only Images
Acquisition #1: water and fat inphase
Acquisition #2: water and fat out of phase
Water image: add acq.#1 + #2
+ =
Called the 2-point Dixon method after the guy who developed it.
Water image
- =
Fat image: subtract acq.#1-acq#2
Fat image
Slide courtesy of Todd Constable, Yale Univ.

## Slide 79: Fat/Water Chemical Shift Effects
Slide courtesy of Todd Constable, Yale Univ.
http://mriquestions.com/chemical-shift-artifact.html

## Slide 80
Low readout bandwidth
High readout bandwidth:
No chemical shift artifact

## Slide 81
Number of voxels shifted depends on fat-water shift and PE bandwidth (Hz/pix)
3.5 ppm x 300 MHz = 1050 Hz fat-water shift
Echo spacing = 0.81ms with R=4 acceleration  1/.20ms = 5 KHz for 180 pix  27 Hz/pix
 3.5ppm x 297 MHz/(27 Hz/pix) = 38 pix shift
Fat saturation in EPI
1.1mm in-plane, 2mm slice
Fat sat OFF
Fat sat ON
Phase encode direction

## Slide 82: Physiological noise in EPI time series: Impact on temporal SNR (μ/σ)
References: Triantafyllou C, et al., NeuroImage 2011
Polimeni J and Wald LL, NeuroImage, 2017
“Noisy tissues”  CSF
Lower noise in white matter
Smaller voxels suffer less partial volume effect from CSF, etc.
(two cases agree more closely)
9x smaller volume!!!

## Slide 83: Quiz!!!

## Slide 84
Solution: Can soften with a little filtering/windowing in frequency domain  causes slight blur
(1) Incomplete fat suppression (chemical shift artifact)
(2) “Zipper” artifact from RF noise contamination
(3) Gibbs ringing from k-space truncation
(4) Head motion artifact
Quiz

## Slide 85
Solution: No easy solution….
(1) Blood flow artifact from hematoma
(2) T2* dephasing from metal object (susceptibility)
(3) RF coil transmit field inhomogeneity
(4) T2* dephasing from sinus cavity air-tissue susceptibility interface
Spin echo
Gradient echo
Quiz

## Slide 86
(1) Blood flow artifact from hematoma
(2) T2* dephasing from metal object (susceptibility)
(3) RF coil transmit field inhomogeneity
(4) T2* dephasing from sinus cavity air-tissue susceptibility interface
Quiz
Solution: Use thinner slices or better B0 shimming
200
-200
0
Hz

## Slide 87
(1) Geometric stretching from phase encode gradient eddy currents
(2) Motion artifact
(3) Fat-water shift along readout direction
(4) No artifact – image shows real anatomical details

## Slide 88
Solution: Reduce echo spacing time, increase iPAT GRAPPA factor, improve B0 shimming
(1) Gradient nonlinearity
(2) Geometric distortion from gradient eddy currents
(3) Motion artifact
(4) Geometric distortion due to B0 offset in body
Quiz

## Slide 90
Cause: Iron oxide suspended in beeswax hair product
McKinstry RC et al., AJR, 2004

## Slide 92: Book recommendations
$35.00 www.lulu.com
(also availible on amazon for 2x this price…)
Very terse, very signal processing oriented. Certainly worth the price.
Wald, 6.S02, 2013 MGH, A.A. Martinos Center

## Slide 93: Book recommendations
$151.00 www.amazon.com
Covers everything about different MRI pulse sequences at a nice level…
Wald, 6.S02, 2013 MGH, A.A. Martinos Center

## Slide 95
Fig. 3: Effects of parallel imaging acceleration on the balance between thermal and physiological noise. Example tSNR maps acquired from 3T fMRI data. With higher acceleration factors (R) the √R and g-factor noise penalty causes the acquisition to be thermal-noise dominated, and so the spatial pattern of tSNR is mainly a function of the distance from the surface coils.
Physiological noise in EPI time series: Impact on temporal SNR (μ/σ)

## Slide 96
ky
kx
180o
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 97
Top 6 features of Fourier transform
2)
  Because we have pretended our object is periodic, we will get a periodic object… (just keep one of them…)
FT-1
Slide courtesy of Lawrence Wald, A. A. Martinos Center, MGH

## Slide 98: Gradient Field Distortions Produce image distortions
Usually correctable in software image recon
Slide courtesy of Todd Constable, Yale Univ.
At Martinos, watch out for gradient nonlinearity in the neck in Bay 5 and Bay 8
On MR/PET scanners gradient nonlinearity is a problem in the arms for attenuation correction

## Slide 99: Field Inhomogeneities: Magnetic Susceptibility Artifacts
Gradient Echo (signal loss & image distortion) Spin echo (image distortion)
Knee joint
Slide courtesy of Todd Constable, Yale Univ.

## Slide 100: Respiratory Gating
Slide courtesy of Todd Constable, Yale Univ.

## Slide 101: Fat/Water Chemical Shift Effects
Dark line (gap) as fat shifted down
Bright line where fat shifted to overlap with tissue
dark line can also form if fat and water out of phase
Slide courtesy of Todd Constable, Yale Univ.

## Slide 102: Body images: fat/water: in- and out- of phase
Fat/water in phase
Fat/water out of phase
Dephasing occurs within a voxel (fat & water in the same voxel)
Slide courtesy of Todd Constable, Yale Univ.